- Ede, 17 January 2020

- Driver programming of Norhof LN2 pumps #915

- For advanced programmers.

- The Norhof pumps work with a processor in the pump.

- All working functions are in the pump itself.

- The Windows software we supply, is used to put in some parameters during setup, and monitor behavior during working.

- For most applications, after that, the software does not need to run and the PC can be disconnected.

- It can be connected again in case of readjustments, or if you have the feeling that something is not working correctly.

- For remote control of switching ON and OFF the pump, there are 4 possibilities, according the manual.

- This is this text in the manual:

# =====================

# Appendix C : Remote control

All the #600, #800 and #900 models can be remotely switched OFF and ON (= in SLEEP or STANDBY), in 4 ways:

1 With a 5 volt signal on pin 4 of the 25p subDconnector, the pump can be switched to SLEEP.

0 Volt = SLEEP, 5 Volt is no-operation.

With a 5 volt signal on pin 5 of the 25p subDconnector, the pump can be switched to STANDBY.

0 Volt = STANDBY, 5 Volt is no-operation.

Connecting these pins to ground (pin 17=18=19=20 ), is also enough to make the signals switch.

So two pushbuttons could do the job also.        SLEEP overrides the STANDBY.

(pull up resistor is 1 kohm,  with 100nF capacitor)

2 With the optocoupler signals (see below)

For programmers:

3  If  NO monitor sofware is standby, so the serial port is free for your own application software, the serial port can be opened at 19200,N,8,1

is Baudrate 19200, No parity, 8 databits, 1 stopbit.

Sending a "pon" command (PumpingON) will set the pump in STANDBY, sending a "pof" command (PumpingOFf) will set the pump to SLEEP.

Sending "rm 19" will give back the pump status register.   result and 2 = pumping, result and 16 = standby,  result and 32 = sleep,  result and 64 = dewar level lower than 5 liters, result and 128 = alarm

4  If the monitorsoftware is running, so the port is occupied, there is a way to tell the monitor software to switch the pump. This is not a very gentle solution, but the only easy way.

The monitor software will check continuously in his working directory (c:\Program Files\NorhofPumpMonitor  as default) for a (empty) file "PON.txt" or "POF.txt"  (or "pon.txt" or "pof.txt" ) .  If found, the pump will switch ON or OFF, and delete the file.

creating a "RM19.txt" file gives back a "RM19.dat" file with the pumpstatus result.

- pin connections 25p subD connector

25p Dconn pins #800 and #900 models, and 600 models from after Jan. 2011:

1=14   AC1    AC or DC 12 to 24 Volts Power supply (min. 4 A.)

13=25  AC2    AC or DC 12 to 24 Volts Power supply (min. 4 A.)

17=18=19=20  system ground

15 opt1C*   optocoupler input 1 neg.

2   opt1A*   optocoupler input 1 pos.  : 0 or 5-24 Volt input.  5-24 V. to switch pump to STANDBY

(use 1 Kohm resistor for 5 Volt, 5K for 24 Volt)

16 opt2C*  optocoupler input 2 neg.

3   opt2A*   optocoupler input 2 pos.  : 0 or 5-24 Volt input.  5-24 V. to switch pump to SLEEP

(use 1 Kohm resistor for 5 Volt, 5K for 24 Volt)

(*) connect C to ground and supply positive signal with resistor to A to switch

(*) OR, connect A to ground to use negative signal on C .

# =====================

- As you can see, RS232 commands can be put directly in the pump, to switch ON and OFF, for all pump models, with the "pon" and "pof" command.

- Only for pump model #915, it is possible to have full control over all functionality, so advanced programmers may write their own program to create a new functionality, if one of the 15 functions does not suit your application.

- The pump work with a driver module as class module for all programming languages, which interacts between your program and the pump.

- This driver takes care of the communication to and from the pump by RS232, and  all raw data from the processor is translated in readable values, taking care of all offsets, calibrations, and non linear characteristics of the used PT100 sensors which are in and on the pump.

- For making your own program, this should be the way how to do that.

- Therefore Windows (any version) should run.

- With the #915 model pump, we supply an example program to set up the driver as a class module, and show how to read and write to and from the driver. This example program is in Visual Basic 6.0. (the first development of this is from 1990).

- Experienced programmers could look into the (short) source code of this, and set this up in Visual C, or C or C++ of whatever programming languages which exists.

- It should not be too much work to get it working as a start.

- Making use of the driver can not harm the pump.

- In case that you can or not want to use Windows, you could try to get things to work making use of the direct RS232 commands, but this is a lot of work, and risky.

- You should read and write directly into the processors memory, to obtain raw data values.

- All data should be converted to real values by yourself, calculating with the correct offsets, calibration values and non-linear  characteristics.

- Also, writing a wrong value on the wrong memory place, could theoretically destroy the pump.

- fe. If by mistake because of writing on a certain place, the heater in the pump goes ON, while the pump is not in the LN2, the heater element will burn out. Also writing on certain places which control the working voltage, could theoretically burn out parts on the main PCB of the pump.

- Mostly writing a wrong value will hang the firmware, causing a reset and all will work again, but fe. writing a wrong value in EEprom  will definitely destroy correct working.

- For if you still want to use your own RS232 connection to a #915 pump, we have a simplified programming manual available, and we could offer some help if you need more.

- But, this is only for the #915 model, and on your request.

- We strongly advice not to do this, and of course this is at your own risk.

- mr. Emile Bisschop, software development & support
LN2 microdosing systems
e-mail direct: dahedi@ziggo.nl
tel.: (00-31)-(0)85-1045715 , workshop Ede
direct: 00-31-6-51577649 , workshop Tienhoven

Main office and workshop:  
NORHOF bv.
Galileilaan 33U
6716BP Ede, Holland

R&D workshop:
NORHOF LN2 Injection Systems bv.
Laan van Niftarlake 44
3612BT Tienhoven, Holland