- Maarssen, 28 december 2011

- Low level programming of Norhof LN2 pumps #915

Reading and writing directly on low level to the RS232 interface is very dangerous.

If written on a wrong place, the processor can act unpredictable and may destroy the pump.

When used low level programming, we can give NO warranty on the pump.

If the pump malfunctions after low level programming, the only solution is to send the pump back to us so that we can load the pump again with the right firmware. But due to wrong programming, also electronic damage may occure.

The Norhof LN2 pumps (upto jan. 2011) contains an Atmel mega8535 Uprocessor, which contains :

8k of flash memory ( for program)

512 bytes of Eeprom (for fixed values)

512 bytes of Sram  (working memory)

In 2011 we started using the Mega16 Uprocessor, which has

16k flash, 512 bytes Eeprom and 1K Sram.

In 2018 we introduced the Atmel Mega32 processor.

During initialising of the processor (power up) all Eeprom values are copied to the relevant Ram spaces in the working memory. During the program, the Eeprom is (almost) never read again.

Uprocessor has a serial interface (RS232) which can read and write values from and to Eeprom and Ram.

For connecting to the interface, the COM parameters are 19200,N,8,1

Command string list:

"i"			returns model and firmware version

"pof"			switches pumping OFF

"pon"			switches pumping ON

"re xxx x"		read Eeprom , address, nr of bytes

"rm xxx xx"		read memRAM, address, nr of bytes

"we xxx xx"		writes Eeprom, address, databyte  	(must be sent double)

"wm xxx xx"		writes memRAM, address, databyte	(must be sent double)

All data and addresses is in hexadecimal values, lower case.

Character send must be terminated with CRLF (0d0a). Characters send to the interface are echoed back, and sends back "ready", or "wrong command". For sending read commands, the commands are single. For write commands, the commands must be send twice (to prevent wrong writing when there is electronic disturbance on the interface cable). Command is only accepted when it was twice the same.

The INFO command

The "i" command gives back what pump it is.

So sending an "i" with CRLF, the pump will return for example:

"    Model:   NORHOF 900 series pump

Version: Ver 7.50

Date:   2 jan. 2012 "

Pumping OFF and ON commands

For switching the pump OFF (= sleep mode) you can use the "pof" command.

For switching the pump ON (= ACTIVE or PUMPING, depending on working mode and sensor(s) status), you can use the "pon" command.

pon and pof commands does not need to be sent double. Single will be accepted.

To read and write Eeprom and Ram memory, example :

Eprom read :

re 010 		= read eeprom &H10, 1 byte    (probably "ed")
re 010 2 	= read eeprom &H10, 2 bytes   (probably "ed 00")
re 010 10 	= read eeprom &H10, 16 bytes (probably "ed 00 87 00 89 00.........)
!!! do NOT write to Eeprom !!!

RAM read and write:

low 32 bytes of Ram is the working registers R0 to R31.

further upto &h5F is I/O registers.

first real RAM byte is &H60

rm 060 		= read memory &H60, 1 byte
rm 060 2 	= read memory &H60 and &H61, 2 bytes
rm 060 10 	= read memory &H60 to &H6F, 16 bytes

On this moment 234 to 23f is not used, so you may test for writing on this place without danger
!!!!! be very sure that you write on an adress were it can do no harm

Example:

"rm 234"  reads  probaly f0 or 0d , because memory is filled with f0 0d before intitialising. If an other value is there, do NOT write here.

"wm 234 3a" = write memory &H234 with value 3a   
writing command must be sent double , so two times directly after each other.

so "wm 234 3a" crlf , "wm 234 3a" crlf

with "rm 234" you can read back if it succeeded, should be now 3a.

If you make use of our driver, the pump913drv.exe, as a class module for VB or C+, the driver itself already sends all write commands double. So then you can send a write command single to the driver.

But when the pump is directly connected to the com port, to your software, the write commands should be send twice, before it is accepted by the pump.

Read temperature and pressure

To read the actual temperature of sensors and pressure in the dewar:

The sensor voltages and the pressure is a value from the ADC, 10bits, so from 0 to 1023 (01ff)

The ADC has 8 channels. The result is copied into mem 80 to 8f, and a 9th value is extraced from a voltage measurement, on 90 and 91
Main sensor is on &H86 and &H87

The pressure sensor is in &H88 and &H89

; AD results space 0x080 (ADCAVEST)
; -----------------------------------------------------------------------
$080 and 081 ; Vessel sensor value pa0 (Low,High)        ' 2x PT100 sensors in serial, in bottom of dewar
$082 and 083; TMB sensor value pa1 (Low,High)           '  ceramic PT100 sensor  in outlet tube in pumphead

$084 and 085; Extra sensor no.1 value pa2 (Low,High)   ' external PT100 sensor "extra"
$086 and 087; Main sensor no.2 value pa3 (Low,High)   ' external PT100 sensor "main"

$088 and 089; Presure sensor value pa4 (Low,High)       ' pressure sensor for pressure in bottom of dewar
$08a and 08b; External flow value pa5 (Low,High)        ' input analogue 0-5Volt for external flow preset

$08c and 08d; Internal flow value pa6 (Low,High)         ' flowset potmeter on pump
$08e and 08f;; Int/ext temp. value pa7 (Low,High)        ' temperature preset potmeter on pump (#910 and #915)"
$090 and 091; Power supply voltage (low,High)            ' measured input power voltage 
 
Example to read the actual pressure:
so the pressure value is in  088 and 089

"rm 088 2" will give 23 00 to 40 00 when the pump is out of the LN2, so when there is NO pressure on the pressure sensor.

The zero value for the pressure sensor is is also needed to calculate the actual pressure.

Each pump is calibrated for this zero pressure value in production.

Zero calibration values are in EEprom on:

Eeprom address                                                                      value probably

$010 and $011;   Vessel sensor cold (-196) 			0xed, 0x00

$012 and $013;    TMB sensor cold (-196)			0x87, 0x00,

$014 and $015;    extra sensor cold (-196)			0x8f, 0x00

$016 and $017;    main sensor cold (-196)			0x8f, 0x00

$018 and $019;    vessel room (-100)   				0xa0, 0x01

$01a and $01b;    TMB heat(+49)				0x0f, 0x01

$01c and $01d;    pressure sensor 0mBar				0x23, 0x00

$01e and $01f;    feeding height compensation			0x0f ,0x00

so fe. for reading the zero calibration for the pressure sensor ,

"re 01c 2" will give the calibrated offset, probably 23 00 ( or 24 00, or 25 00, or 2a 00), somewhere between 23 and 40 
(value can be different after calibration)

BUT,  reading Eeprom is not a good habit, because it will use interupts and halt the program shortly while reading.

Most zero calibration values are copied from Eeprom to RAM during initialising.

So it is much better to read this zero calibration value from RAM.

The zero calibration for the pressure sensor is on RAM 0c8 and 0c9 and always accessible.

Most other zero calibration values (such as for the sensors) should be read out of Eeprom once, and stored in your own program to use them.

Reading pressure:
The actual pressure is "rm 088 2" minus "rm 0c8 2" times the scale.
The scale is 0.54288 mBar per step,

so pressure in dewar is ("rm 088 2") - ( "rm 0c8 2" ) * 0.54288 mBar
Note that data results are in hexadecimal values, and should be converted to decimal first.

If you have a 35 liter dewar, and the dewar is completely filled, reading the pressure while the pump is in SLEEP, would give a value corresponding the dewar LN2 level.

A full 35 liter dewar would give decimal 52 (hex 34)

The 35 liter dewar can hold 36 cm of LN2 upto the neck. (the 50 liter dewar = 42 cm)

The pressure tube starts 10mm from the bottom of the dewar.

The pressure of 35cm LN2 in the tube = 35 * .808 =  28.28 mBar.

28.28 mBar / 0.542888 = 52,09 steps from the ADC.

So when the pressure value ("rm 088 2") - ( "re 0c8 2" )  is 52, the dewar is full.

When the pressure value ("rm 088 2") - ( "re 0c8 2" ) is 0, the dewar is empty (10 mm left)

When the value is 26, the dewar is 50% filled.

Reading dewar LN2 level:

In the pump, the pressure value while not pumping is stored in RAM, to correspond with the dewar LN2 level. When in SLEEP, this value is updated. When in PUMPING, this value is decreased a little depending of the heater actions. So this value is decreased theoreticaly during pumping, but actualised when in SLEEP again.

The value for the dewar level is stored in RAM on address 0ce and 0cf, but the zero calibration value should be substracted to become the actual level value in ADC steps

RAM address LN2 level = 0ce and 0cf

So the dewar level can be read as ("rm 0ce 2") - ( "re 0c8 2" ), where a value of decimal 52 means a full 35L dewar, and a value of 0 means an empty dewar (10 mm left).

External sensors (MAIN and EXTRA sensor, marked as ** and * on the pump.

In most working modes the MAIN sensor is used. In some modes also the EXTRA sensor is used.

the external sensors are
.equ ADCEX1 =  $084 ; External "extra" sensor  value pa2 (Low,High)
.equ ADCEX2 =  $086 ; External "main" sensor  value pa3 (Low,High)

For knowing the exact temperature you need to know  the "zero" (calibration -196) and  "warm" (calibration +30)
zero calibration is in table above,

Warm values are in Eeprom, if used .

Calibration values depend on how the sensor is used, on 33 mA current (selfheating, for level detection), or on 1 mA current (for temperature measurement)

Warm calibration values are in Eeprom on:

address                      					 value probably 
$090 and $091; vessel empty sensor  30 C                	0x7d, 0x02

$091 and $092; TMB  sensor                  	  	  	0x50, 0x01

$093 and $094; 2nd sensor   extra   +30C         	  	0xdd, 0x01

$095 and $096; 1nd sensor   main   +30C     	  	0xdd, 0x01

$097 and $098; Vess room (not used)  		  	0x00, 0x00

$099 and $099; TMB heat  (not used)  		  	0x00, 0x00

$09a and $09b; PreWA prepressure   (not used)	  	0x00, 0x00

$09c and $09d; Precom3		   (not used)	0x00, 0x00

Note that only for the PT100 sensors this value is present

the actual temperature of extra sensor 1 is (rm 084 2) - ( re 014 2)
the actual temperature of main sensor 2 is (rm 086 2)- ( re 016 2)
but remember to read the Eeprom values only at initialising your program

this is in steps of the ADC., from -196 degrees.

The ADC steps is a more or less logaritmic scale corresponding to the PT100
characteristics ( see PT100 specs)

To know the exact value in degrees C, you must know the +30 degrees point in the scale
 
Calculating into degrees is done in normaly the driver.

The PT100 sensors are connected with a resistor of 135 Ohm to the +5 Volt for the 33 mA sensor,

or 4K7 to +5 Volt for the 1 Ma sensors.

With the characteristics of the PT100, a tabel could be created with voltages belonging to the temperatures. Comparing the ADC value to this table gives the exact temperature.

Pump Working states

The pump knows 3 working states.  SLEEP, or ACTIVE (standby), or PUMPING .

In general, the SLEEP mode does not pump, but keeps the heater in the pumpflowtube (the TMB) ON to keep it on 50 degrees. In SLEEP all alarms are reset. In SLEEP is the degass solenoid valve always open.

The ACTIVE state lets the pump PUMPING or NOT PUMPING, depending on sensor(s) and/or timer.

The PUMPING bit then is ON or OFF in ACTIVE state.

Pumping mode SLEEP, or ACTIVE, or PUMPING is in the
pumpstatus register

; Pump status register (R25) = RAM address &H19
; -----------------------------------------------------------------------
.def PStatus = R25 	; Pump Status register

.equ rs_recv =    bit0 	;=1 ; 1 RS232 character received
.equ Pumping = bit1 	;=2 ; 2 Pump request
.equ Valve_on = 2 	;=4 ; 4 Valve closed
.equ TMB_on = 3 	;=8 ; 8 TMB in heating mode
.equ Active = 4 		;=16 ; 16 Pump status : active
.equ Sleeep = 5 		;=32 ; 32 Pump status : sleep
.equ Warn_on = 6 	;=64 ; 64 Pump status : warning
.equ Alarm_on = 7 	;=128 ; 128 pump status : alarm

memory place is &H19, so "rm 19" gives this status ,

Analysing bit 1, bit 4 and bit 5 gives the pump status.

The other bits can vary.

Here is the possibility to check in what status is the pump.

This is NOT the place to change status.

Before the status is changed , some routines must run to set and/or reset some values for this status.

When simply only here the status is changed, perhaps the heater could still be ON while the (new) status is SLEEP.

So for changing the status use the RS232 register RSConSet. (in #915 effective only), or the pon and pof commands.

Also here you can read the status for the alarms, and the Warning.
On an ALARM, the pump stops and beeps

On the WARING, all pumping continues, but it will beep every 30 seconds because the dewar is almost empty (<5 liters LN2 left)

Alarm register

If  there is an alarm, the alarm register is
RMSTRT = 060
.equ Alarm = RMSTRT + 2 	; Alarm Flag register   ( so on &H62)
  .equ HOT_AL_ON = bit0 	; Vessel to warm (ROOM TEMP) ,dewar empty
  .equ VES_AL         = bit1	; Vessel sensor alarm flag (ALARM_CHECK)
  .equ TMB_AL = 2 		; TMB sensor broken flag (ALARM_CHECK)
  .equ EX1_AL = 3 		; extra sensor broken flag (ALARM_CHECK)
  .equ EX2_AL = 4 		; main sensor broken flag (ALARM_CHECK)
  .equ FROZ_AL = 5 		; rise pipe frozen alarm
  ;equ   = 6 			; container warmer than -140 (#605 pump only)
  ;equ   = 7 			; 605 filling too long              (#605 pump only)

the alarm mask depends on in what mode is the pump, and how many sensors
there are present.

Alarm mask register

.equ AlarmMask = RMSTRT + 3; (&D063)  Alarm Mask Flag register

.equ  VES_AL_RO  = bit0	;=1   Enable Vessel Room temp check -100

.equ VES_AL_MA  = bit1	;=2   Enable Vessel alarm check     broken

.equ TMB_AL_MA  = 2	;=4   Enable TMB alarm check

.equ EX1_AL_MA  = 3	;=8   Enable extra EX1 alarm check

.equ EX2_AL_MA  = 4	;=10  Enable main EX2 alarm check

.equ FROZ_AL_MA = 5      	;=20  rise pipe frozen alarm

; 6     				;=40  container warmer than -140 (605)

;equ FILL_AL_MA =7      	;=80  605 filling too long 7-12-08

so if (rm 062 )AND( rm 063 ) is >0 then there is an alarm.

 
RS232 commands registers

Switching the pump in the different STATES

With the RS232 registers the pump can be controlled by RS232

; RS232 related registers to allow controlling the pump by RS232

;				0x110

; -----------------------------------------------------------------------

.equ RSTmpSetL = RSREG + $0	;(&h110) RS232 temperature setpoint	    (LSB)

.equ RSTmpSetH = RSREG + $1	; 				   	 (MSB)

.equ RSFlwSetL = RSREG + $2	;(&h112) RS232 flow setpoint		    (LSB)

.equ RSFlwSetH = RSREG + $3	;				    	(MSB)

.equ RSConSet	 = RSREG + $4		;(&h114) RS232 pump on/off control

.equ  RSCSACT	= bit0		;=1 RS232 control active bit

.equ  RSCSPUMP	= bit1		;=2 RS232 control Pumping bit

.equ  RSCSEXAL	= 2		;=4 RS232 control Extra sensor alarm mask

.equ  RSCSMAAL 	= 3		;=8 RS232 control Main sensor alarm mask

.equ  RSCSEXLR	= 4		;=10 RS232 control Extra sensor led red

.equ  RSCSEXLG	= 5		;=20 RS232 control Extra sensor led green

.equ  RSCSMALR	= 6		;=40 RS232 control Main sensor led red

.equ  RSCSMALG	= 7		;=80 RS232 control Main sensor led green

;tm 114

in the RSConSet the pumping state can be set, with the alarm mask for the two external sensors.

For mode E, the leds on the pump will operate according the mode.

For mode F, the leds on the pump can be switched ON and OFF.

RSTmpSetL can be set for mode E (temperature control mode)

RSFlwSetL can be set for both mode E and F.

Of course there is a lot more details, but most can be seen in the example VB programs.

Contact us for more specific instructions you may need.

NB: Using the RS232 interface causes the processor to need time for the communication.

If reading is continuous, the processor can HALT and so proper working can be disturbed.

Reading parameters from the pump should always be done with a certain interval in time, so that the processor has enough time to finish its other process.

So, read unimportant parameters once every 5 seconds, and minimize reading fast needed parameters, and read these parameters fe. every 0,3 seconds, but not faster.